# YT-98H Gas Transmitter

Reads 5 gas sensors over Modbus RTU (RS-485). Works on Windows, Linux and macOS.

## Start here

chmod +x ./dev_path.sh

chmod 666 [/dev/USB0 OR WHATEVER DEVICE IT IS] # use this to map the device (so it will change when connected directly)

uv run yt98h_modbus.py detect     # find the device

uv run yt98h_modbus.py map        # see every register and every gas value

uv run yt98h_modbus.py watch      # live values, updates every second



uv run installs everything by itself. No Python setup needed.

To make the CSV file for the vendor, run this and press Ctrl+C after a few seconds:

uv run modbus_register_sweep.py



It writes sweep_log_<date>_<time>.csv.

## How to talk to the device

speed      9600 8N1

function   FC03  (read holding registers)

addresses  1, 2, 3, 4, 5

read       start = 0, count = 32



Always read all 32 registers in one request. If you read one register alone, this device gives back the wrong data.

## The 5 sensors

Each sensor has its own Modbus address.

| Address | Gas | Unit | Full range | REG_31 raw | Real value |
| --- | --- | --- | --- | --- | --- |
| 1 | H2S | ppm | 200 | 0 | 0 ppm |
| 2 | CO | ppm | 2000 | 0 | 0 ppm |
| 3 | O2 | %VOL | 30 | 209 | 20.9 %VOL |
| 4 | LEL | %LEL | 100 | 0 | 0 %LEL |
| 5 | CO2 | ppm | 50000 | 834 | 834 ppm |

Zero is correct in clean air. Only O2 and CO2 show a value in a normal room.

## The 32 registers

One full read. Same data you get from uv run yt98h_modbus.py map.

| Reg | Addr 1 | Addr 2 | Addr 3 | Addr 4 | Addr 5 | What it is |
| --- | --- | --- | --- | --- | --- | --- |
| r0 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r1 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r2 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r3 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r4 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r5 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r6 | 201 | 201 | 201 | 201 | 201 | same on all sensors |
| r7 | 201 | 201 | 201 | 201 | 201 | same on all sensors |
| r8 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r9 | 10135 | 10135 | 10135 | 10135 | 10135 | air pressure, 1013.5 hPa |
| r10 | 201 | 201 | 201 | 201 | 201 | same on all sensors |
| r11 | 180 | 180 | 180 | 180 | 180 | same on all sensors |
| r12 | 1 | 1 | 1 | 1 | 1 | same on all sensors |
| r13 | 1 | 1 | 1 | 1 | 1 | same on all sensors |
| r14 | 1 | 1 | 1 | 1 | 1 | same on all sensors |
| r15 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r16 | 31 | 31 | 31 | 31 | 31 | same on all sensors |
| r17 | 1 | 1 | 1 | 1 | 1 | same on all sensors |
| r18 | 1 | 1 | 1 | 1 | 1 | same on all sensors |
| r19 | 3 | 3 | 3 | 3 | 3 | same on all sensors |
| r20 | 64113 | 64113 | 64113 | 64113 | 64113 | same on all sensors |
| r21 | 5 | 5 | 5 | 5 | 5 | how many sensors, 5 |
| r22 | 1688 | 1688 | 1688 | 1688 | 1688 | device ID |
| r23 | 59727 | 59727 | 59727 | 59727 | 59727 | device ID |
| r24 | 12802 | 12802 | 12802 | 12802 | 12802 | device ID |
| r25 | 2 | 1 | 3 | 65 | 18 | which gas |
| r26 | 0 | 0 | 2 | 3 | 0 | unit |
| r27 | 1 | 0 | 1 | 0 | 0 | decimal places |
| r28 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r29 | 2000 | 2000 | 300 | 100 | 50000 | full range |
| r30 | 0 | 0 | 0 | 0 | 0 | same on all sensors |
| r31 | 0 | 0 | 209 | 0 | 834 | THE GAS VALUE |

r0 to r24 are the same on every sensor. They are device settings. r25 to r31 change per sensor. Those are the ones you want.

## How to get the gas value

value = REG_31 / 10 ** REG_27



| Register | Meaning |
| --- | --- |
| REG_25 | which gas: 1=CO, 2=H2S, 3=O2, 18=CO2, 65=LEL |
| REG_26 | unit: 0=ppm, 1=mg/m3, 2=%VOL, 3=%LEL |
| REG_27 | decimal places |
| REG_29 | full range |
| REG_31 | the gas value |

Example, address 3: REG_31 = 209, REG_27 = 1, so 209 / 10 = 20.9. REG_26 = 2, which means %VOL. Answer: 20.9 %VOL oxygen.

Example, address 5: REG_31 = 834, REG_27 = 0, so no change. REG_26 = 0, which means ppm. Answer: 834 ppm CO2.

## Change settings

Open yt98h_modbus.py. The top has a block marked CONFIG. EDIT THIS PART. Most people never need to touch it.

| Setting | What it is |
| --- | --- |
| PORT | Leave "" and it finds the port by itself. Or write "COM5". |
| BAUD | Speed. 9600 for this device. |
| ADDRESSES | The sensor list: [1, 2, 3, 4, 5] |
| GAS_TYPES | Which number means which gas |

If a sensor shows as code 7 instead of a name, add one line to GAS_TYPES: 7: "NH3",

If the script picks the wrong port, pass it on the command line instead:

| Your computer | Add this |
| --- | --- |
| Windows | --port COM5 |
| Linux | --port /dev/ttyUSB0 |
| macOS | --port /dev/cu.usbserial-XXXXXXXX |

## The files

| File | What it does |
| --- | --- |
| yt98h_modbus.py | Main tool. detect, map, watch. |
| modbus_register_sweep.py | Makes the CSV for the vendor. |
| Get-ModbusDiagnostics-1.ps1 | Windows only. Checks your PC, not the sensor. |

## If it does not work

- Swap the A and B wires. This is the answer most of the time.

- Check 24V power is on and the screen is lit.

- Wait 2 to 5 minutes. Sensors need to warm up.

- Set Output Mode to RS485 in the screen menu.

- Close other programs using the port (Modbus Poll, PuTTY).

- Run uv run yt98h_modbus.py detect. It tries every speed and address.

Two common mistakes:

- pymodbus 3.14 uses device_id=, not slave=. Old code crashes.

- On macOS use /dev/cu.*, never /dev/tty.*. It will freeze.

CO2 updates slowly, every 15 to 45 seconds. Wait a full minute before you think it is broken.